## Slide 1: Li7200 Obsolescence
Li7200 vs Li7200 RS

## Slide 2: Background
The Li7200 is currently being phased out by the latest Licor model Li7200 RS.
Testing was conducted to verify:
  Connectivity to the CI pipeline
  If the Li7200 RS was comparable to the Li7200
  If the Li7200 RS passed requirements?

## Slide 3: Test Design
Li7200 and Li7200 RS were installed in parallel with their own closed path vacuum systems.
Test was broken out into two phases.
  Phase 1: 2021-04-22 to 2021-05-01
  Phase 2: 2021-05-07 to 2021-05-16
After SAE had 7 days of quality data the Li7200s were switched such that if there were any sampling biases introduced by the different closed path vacuum systems, they would be visible in the data.

## Slide 4: QFQM
Many flags were raised on several days and are highlighted in the final plots
Flagged data comes out as NaN so removing that day’s data did not heavily impact the statistics/plots

## Slide 5: Executive Summary
Short discussion about the results and statistics

## Slide 6: Executive Summary (cont.)
CO2 Fluxes
H2O Fluxes

## Slide 7: Executive Summary (cont.)
CO2 Fluxes
H2O Fluxes
CO2 Fluxes
H2O Fluxes

## Slide 8: Executive Summary (cont.)

## Slide 9: Additional Plots – CO2 1-min – Raw Values

## Slide 10: Additional Plots – CO2 1-min – Daily Box Plot

## Slide 11: Additional Plots – CO2 1-min – Difference

## Slide 12: Additional Plots – CO2 1-min – Li7200 ~ Li7200RS

## Slide 13: Additional Plots – H2O 1-min – Raw Values

## Slide 14: Additional Plots – H2O 1-min – Daily Box Plot

## Slide 15: Additional Plots – H2O 1-min – Difference

## Slide 16: Additional Plots – H2O 1-min – Li7200 ~ Li7200RS

## Slide 17: Additional Plots – frt00 1-min – Raw Values

## Slide 18: Additional Plots – pdiff 1-min – Raw Values

## Slide 19: Results